Appendix B: External computer data formats (air-cooled version)

Rapid Scan Parameters to be transmitted from a control computer to the scan coil driver.

Units, Ranges and Resolutions for air-cooled version.

All parameters are 12-bit resolution:

| PARAMETER | UNITS | RANGE (Numerical) | RANGE (Normal Usage) | RESOLUTION 1 LSB = |
| --- | --- | --- | --- | --- |
| Scan Width | Gauss (p-p) | 0 to 20.475 | 0.5 to 10 | 0.005 Gauss |
| Scan Frequency | Hertz | 500 to 4,595 | 1000 to 4,500 | 1 Hertz |
| SpecJet Trigger Phase | Degrees | 0 to 409.5 | 0 to 359.9 | 0.1 Degree |

RS-232 Data Transmission Format

Data are transmitted on a 9-pin or 25-pin RS-232 serial port with the following format:

- A quiet interval timeout of 1 second minimum must precede the block transmission. This can be most efficiently implemented by starting a one second timer at the end of each block and checking to see that the timer has expired when a new transmission is required. In normal operation it will have long since expired and there will be no artificially imposed delay. No delay is required between the words of a block transmission.

- A transmission consists of 6 words. Each RS-232 word contains one 8-bit byte of data plus the normal 1 start bit and 1 stop bit..

- For each parameter the most significant byte is first, followed by the least significant byte. Within each byte the most significant bit is first.

- If one parameter changes, all 3 must be retransmitted in a block.

- The RS-232 format is to be 1 stop bit, 8 data bits, no parity, 9600 baud.

Example of a transmission in which the following values are transmitted:

Scan Width = 3.00 Gauss

Scan Frequency = 3,000 Hz

Digitizer Trigger Phase = 25 degrees

| Word number (transmitted in this order) | Description | Bit pattern MSB … LSB |
| --- | --- | --- |
| 1 | Sweep Width High Byte | 00000010 |
| 2 | Sweep Width Low Byte | 01011000 |
| 3 | Sweep Frequency High Byte | 00001001 |
| 4 | Sweep Frequency Low Byte | 11000100 |
| 5 | Digitizer Trigger Phase High Byte | 00000000 |
| 6 | Digitizer Trigger Phase Low Byte | 11111010 |